## Slide 1: Hearing preservation cochlear implant acute and chronic trials
June 2019
Yoojin Chung

## Slide 2: List of all hearing preservation experiments
C60 – acute intra-cochlear 4/5/2018
  No resp after CI
C63 – acute intra-cochlear 4/27/2018
  20 dB ↑ bullotomy, no acoustic response post insertion
C66 – acute extra- & intra-cochlear 5/18/2018
  10 dB ↑ bullotomy, 10 dB ↑ cochleostomy, 40 dB ↑ opening cochlestomy
C68 – chronic extra-cochlear 6/4/2018
  No change post implantation
C72 – chronic extra-cochlear 7/16/2018
  ~30 dB shift post implantation (immediately).
C73 – chronic extra-cochlear 7/31/2018
  No change post implantation
C68 – acute intra-cochlear 4/3/2019
  No change bullotomy, 10 dB ↑ RW drilling, 20 dB ↑ opening RW membrane and CI insertion, 20 dB ↑ RW packing (also perilymph leaking)
C90 – chronic intra-cochlear 4/23/2019
C76 – chronic extra-cochlear 4/30/2019
C89 – chronic intra-cochlear 5/14/2019
C73 – acute extra- & intra-cochlear 5/30/2019

## Slide 3: C60 Acute experiment
April 5, 2018, Left ear
This was the first acute experiment. We tried using closed acoustic delivery system and had a lot of difficulties.
Right ear – could not get a good calibration
Left ear – ABR done before surgery and after electrode insertion

## Slide 4
Before
After
C60 acute experiment
Click ABRs before surgery and after CI insertion
No acoustic response post insertion

## Slide 5: C63 Acute experiment
April 27, 2018, Left ear
This was the first acute experiment. We tried using closed acoustic delivery system and had a lot of difficulties.
Right ear – TM perforation
Left ear – ABR done before surgery, after bullotomy, and after intra-cochlear electrode inserted

## Slide 6
C63 acute experiment
Click ABRs before surgery, after bullotomy, and after insertion
20 dB ↑ bullotomy, no acoustic response post insertion
Bulla opened
Before surgery
After insertion 1
After insertion 2

## Slide 7: C66 Acute experiment
May 18, 2018, Left ear
First bullotomy made – here there was minimal elevation following drilling. Did Woongsang do something differently?
Place RW implantation
Made a partial cochleostomy – left a thin layer of bone and membrane
Cochleostomy punctured
Note is incomplete – it appears that we did not try inserting intracochelar electrodes.

## Slide 8
Bulla exposed
Bulla opened
RW electrode placed
Partial
cochleostomy made
Cochleostomy punctured
C66 acute experiment
Click ABRs at each step
10 dB ↑ bullotomy, 10 dB ↑ cochleostomy, 40 dB ↑ opening cochlestomy

## Slide 9: Animal IC: C68
Round window implantation in L ear
June 4, 2019
ABRs done immediately after surgery
Left (implanted)
Right

## Slide 10: Animal IC: C72
Round window implantation in L ear
July 16, 2018
ABRs done before and immediately after surgery
~30 dB elevation in the implanted ear
Left (before)
Left (after)
Right (before)

## Slide 11: Animal ID: C73
Animal history: RW implanted in Left ear July 30, 2018.
  Implant: HL-8, most apical electrode placed on RW.
  No elevation in click ABR threshold post RW implantation. (Tested immediately after surgery.)
Acute extra & intra-cochlear implantation experiment on May 30, 2019.

## Slide 12
Before surgery
Bulla opened
RW partially drilled
/Applied Healon
CI inserted
C68 acute experiment
No change bullotomy, 10 dB ↑ RW drilling, 20 dB ↑ opening RW membrane and CI insertion, 20 dB ↑ RW packing (also perilymph leaking)
RW packed

## Slide 13: C73 Acute experiment
May 30, 2019
In the Right ear
First placed RW implantation (modified hardball electrode from Cochlear)
Opened RW and inserted HL-8
Measured acoustically and electrically evoked ABRs throughout (all in R ear unless specified otherwise)
Animal perfused with 4% PFA in PBS and cochleas harvested

## Slide 14
C73 Acute experiment
Tone-pip ABR before surgery

## Slide 15
C73 Acute experiment
After opening the bulla
30-40 dB elevation
Bulla was opened with a diamond burr

## Slide 16
C73 Acute experiment
Acoustic ABR after placing the RW implant
Amplitude seems wrong. I might have changed it to x500. (Note indicates it was changed after this measurement).

## Slide 17
C73 Acute experiment
EABR with RW stimulation

## Slide 18
C73 Acute experiment
Acoustic ABR after removing the round window niche and drilling the RW edge.
There was some leak.
~20 dB elevation

## Slide 19
C73 Acute experiment
Tone-pip ABR after HL-8 insertion
RW is not plugged
There was a lot of liquid at the end – completely filling the middle ear.

## Slide 20
C73 Acute experiment
Acoustic ABR after inserting HL-8 (bulla & RW opened)
  Thd = ~-30 dB
Acoustic ABR after plugging the RW and closing the bulla
Electrical ABR with HL-8
Next page: Tone-pip

## Slide 22: C73 acute experiment summary
30-40 dB elevation in acoustic threshold immediately after opening the bulla
  Acoustic thd increases with bulla opening
  This elevation is probably due to the drilling noise
  There was no elevation in click ABR following RW implantation in the R ear
Some hearing preservation following CI insertion
Overall there are multiple factors:
  Thd shift due to drilling – this might be actively recovering during the exp
  Thd shift due to the bulla opening – might have a confounding effect
  Perilymph leakage
  Closing the RW to stop the leakage also has effect in the high frequency

## Slide 23: Animal IC: C90
Hearing preservation CI implanted in R ear Apr 23, 2019
  Cochlear HL-8
Acoustic ABR on May 17, 2019
  Some resp at high frequencies

## Slide 24
C90 pre-op ABR
Click & tone-pip ABR
Apr 20, 2019

## Slide 25
C90 post-op ABR
Click ABR and electrical ABR
Apr 30, 2019

## Slide 26
C90 post-op ABR
Click & tone-pip ABR
May 17, 2019

## Slide 27
C90 post-op ABR
Click & tone-pip ABR
May 24, 2019

## Slide 28: Animal IC: C89
Hearing preservation CI implanted in R ear May 14, 2019
  Cochlear HL-8
Acoustic ABR on May 24, 2019
  No resp.
  Repeat again in a few weeks